## Slide 1: MTEC
Group 9
Francis Bato
Bishoy Botros
Erich Dondyk
Nghia Matt Nguyen

## Slide 2: What is the MTEC?
MTEC
  Material Testing Equipment Controller
Application
  Material Testing for orthopedics of war veterans’
Sponsor
  Dr. Gordon & MMAE

## Slide 3: Goals & Objectives
Motor control to simulate the stepping motion of a human foot
User-friendly interface
Robust Hardware
Standalone device with data storage
Real-time data display
Compatible with Windows, Mac
2 Modes
  Actuator/Load Cell (AL Mode)
    Motor Control + Data Acquisition
  Load Cell / Transducer (LT Mode)
    Data Acquisition only

## Slide 4: How does it work ?
GUI
MTEC

## Slide 5: Specifications & Requirements
Dimensions
  200.7 x 279.4 x 76.2 mm
Weight
  < 5 lbs.
Microcontroller
  > 20 I/O, > 10 ADC, 4 I/O ADC, 8 PWM
Operating voltage: 3.3V – 15V
Motor
  Single actuator fits within a 1-1.5 in2 area
  Applied force of 25 lbs.
Sensors
  Up to 8 load cells
    Sustains 50 lbs. each
  2 displacement sensors (transducer)
    Threshold of 10-20 mm

## Slide 6: MTEC Block Diagram

## Slide 7: MTEC Software Diagram

## Slide 8: Motor Control

## Slide 9: Motor Control
AC vs. DC vs. Servo vs. Stepper motors
Bidirectional motion
Speed control
Motion control ( pushing or pulling on the material)
Pulse Width Modulation for motion control

## Slide 10: H-bridge
Allows for switching the voltage input for bidirectional movement.
CCP Vs. ECCP pins

## Slide 11: Linear Actuator
Model L12-50-210-06-I
50 is stroke length in mm
210 is gear ratio giving up to 150 N ≈ 33lb
06 is voltage
I is for microcontroller interface.
30 cm

## Slide 12: Power Supply

## Slide 13: Power Distribution

## Slide 14: Comparison
Linear Regulator | Switching Regulator
Excess voltage must be dissipated (Heat) | Efficient in conversion of electrical power (less heat)
Easier integration | Complex circuit integration
Inexpensive | A bit costly
Less efficiency | Much more efficient
Size and weight issues | Smaller size and lighter weight
LM 7805, 06 | LM2598, LM 2599 (error flag), LM2673 ( adjustable current limit) Output Range 1.23 – 37 volts

## Slide 15: Sensors

## Slide 16: Displacement Transducer
LD 621 model.
Input DC voltage between 10-30 V at 100 mA.
Output 0 – 10 VDC
Linear relationship between voltage and displacement in mm.

## Slide 17: Load Cell
LCM 300
Rated Output: 2mV/V
Safe Overload: 150% of R.O.
Zero Balance: +/- 3% of R.O.
Excitation (VDC or VAC): 15 Max
Bridge Resistance: 700 ohms
Calibration Test Excitation: 10 VDC
Capacity: 250 lbs / 1112 N

## Slide 18: Wheatstone Bridge I
Load cell consists of a Wheatstone bridge circuit. 2 corners are used for voltage supply and 2 are output signal.
Voltage supplied in excitation will be 10V.

## Slide 19: Wheatstone Bridge II
Load cell output is 2mv/V. With 10V excitation, the load cell output signal will be 20mV.
+ Output (Tension)
- Output (Compression)

## Slide 20: Op Amp
Needs to be amplified to about 5V for the microcontroller. 4096mV/20mV ~ 205x gain.
Load cell output signals connect to op amp for gain before being connected to A/D pin

## Slide 21: Microcontroller

## Slide 22
Microcontroller Selection
Decided 8-bit technology.
- Fit for purpose
- Simplicity
Originally intended to use a PIC18F4550.
Due to requirement alterations, a MCU with 8 PWMs was necessary.
Only two PIC18 families met these requirements.
Decided PIC Microcontroller.
- Wide array of options
- Performance
- Programmable in C

## Slide 23
Microcontroller Selection
PIC18FXXK22 & PIC18FXXK90
- Identical in most aspects
- Package: TQFP (surface mounted)
PIC18FXXK90 has display controller incorporated.
- Unnecessary feature.
Selected most powerful version of the PIC18F87K22.
Device | Price
PIC18F65K22 | $2.39
PIC18F66K22 | DISCONTINUED
PIC18F67K22 | DISCONTINUED
PIC18F85K22 | $2.66
PIC18F86K22 | $2.97
PIC18F87K22 | $3.21

## Slide 24
MCU Development Board
To program a surface mounted MCU it must be mounted on PCB with a ICSP circuit incorporated.
For prototyping purposes a development kit will be used.
PIC18 Development Kit. $165.00

## Slide 25: Graphical LCD

## Slide 26: Graphical LCD Display Module
Model: CFAG240128L-TMI-TZTS
Manufacturer: Crystalfontz
Specifications:
Graphical LCD Display
240x128 Resolution
White Edge LED Backlight
STN Negative, Blue
Negative Voltage Generator
4-wire Resistive Touch Screen

## Slide 27
Pin | Symbol | Type | Description | Specifications
1 | FG | Ground | Frame ground | 0V
2 | Vss | Ground | Ground | 0V
3 | Vdd | Power | Power supply. +5V | +5V
4 | V0 | Power | LCD contrast | V0= -8.1V for initial setting
5 | WR | Control Line | Data write | WR = L
6 | RD | Control Line | Data read | RD = L
7 | CE | Control Line | Chip Enable | CE = L
8 | C/D | Control Line |  | Command write: WR=L , C/D=H Data write: C/D=L Status read: RD=L, C/D=H Data read: C/D=L
9 | Vee | Power | Negative voltage output | -22V
10 | RESET | Control Line | Resets module | Normal = H ; Initialize T6963C = L
11 | DB0 | Data Line | Data bus | LSB
12 | DB1 | Data Line | Data bus | 
13 | DB2 | Data Line | Data bus | 
14 | DB3 | Data Line | Data bus | 
15 | DB4 | Data Line | Data bus | 
16 | DB5 | Data Line | Data bus | 
17 | DB6 | Data Line | Data bus | 
18 | DB7 | Data Line | Data bus | MSB
19 | FS | Control Line | Font select | 6 * 8 = H ; 8 * 8 = L
20 | RV | Control Line | Reverse | Reverse = H ; Normal = L
Electrical Requirements
Supply Voltage = +5V
Input High Voltage = +2.8V to +5V
Input Low Voltage = 0V to +0.8V
Supply Current = 28.2mA (typical)
The CFAG240128L display module comes with a Toshiba T6963C display controller.
The T6968C has become an industry standard among small sized display modules.
Contrast Control
(Requires a negative voltage)
GLCD Display Controller

## Slide 28
Pin | Description | Specifications
1 | X1 | Digital/analog-to-digital pin
2 | Y1 | Digital/analog-to-digital pin
3 | X2 | Digital/analog-to-digital pin
4 | Y2 | Digital/analog-to-digital pin
The CFAG240128L display module has a 4 wire resistive touch screen.
- Durable, 5 million touches.
- Simple MCU integration.
- Enhances user interface.
The touch screen connects to digital/analog to digital pins.
Resistive Touch Screen

## Slide 29
GLCD |  | MCU | 
Pin | Symbol | Pin | Symbol
5 | WR | 54 | RB4/KBI0
6 | RD | 53 | RB5/KBI1/T3CKI/T1G
7 | CE | 52 | RB6/KBI2/PGC
8 | C/D | 47 | RB7/KBI3/PGD
10 | RESET | 46 | RC5/SDO1
11 | DB0 | 72 | RD0/CTPLS
12 | DB1 | 69 | RD1/T5CKI/T7G
13 | DB2 | 68 | RD2/PSP2/AD2
14 | DB3 | 67 | RD3/PSP3/AD3
15 | DB4 | 66 | RD4/SDO2/PSP4/AD4
16 | DB5 | 65 | RD5/SDI2/SDA2/PSP5/AD5
17 | DB6 | 64 | RD6/SCK2/SCL2/PSP6/AD6
18 | DB7 | 63 | RD7/SS2/PSP7/AD7
19 | FS | 62 | RJ0
20 | RV | 61 | RJ1/ALE
Touch Screen |  | MCU | 
Pin | Symbol | Pin | Symbol
1 | X1 | 79 | RH0/AN23
2 | Y1 | 80 | RH1/AN22
3 | X2 | 1 | RH2/AN21
4 | Y2 | 2 | RH3/AN20
GLCD/MCU Interface
The Graphical LCD display requires 15 digital pins.
The touch screen requires 4 digital/analog-to-digital pins.

## Slide 30
GLCD/MCU Schematic

## Slide 31: Data Output

## Slide 32: Data Output
Goal: Provide the user flexibility in performing data logging activities of extensive material testing through the use of multiple, reliable and portable output peripherals.
Master Synchronous Serial Port (MSSP)
  2 Modes: SPI and I2C
Devices to consider:
  Flash Memory
  Universal Serial Bus
  Wi-Fi

## Slide 33: SPI
Designed for single Master-Slave protocol but can be used with multiple slave devices.
High throughput
Supports full duplex
No message limit
Supports higher data rates
More difficult to implement multiple slave systems because of no device addressing
Lower power requirements

## Slide 34: SPI & Slave Operation I
2 Different methods of implementing slave devices
  Chip Select Method
    Parallel configuration with independent slaves
    Control each slave device through chip select
    When slave is disabled, slave
    goes into a high impedence
    state that does not interfere
    with active slave and ignores
    data sent

## Slide 35: SPI & Slave Operation II
  Daisy Chain Method
    Data is cascaded through all the slave devices
    Requires clock compatibility and same bit configuration among all slave devices
    Clock polarity must be checked in order to determine edges of clock signal on which the data
    is driven and sampled
    Software implementation
    heavy

## Slide 36: MSSP: SPI
SPI using Slave Select was chosen
  Familiarity
  Ease of implementation
  High throughput
  Although I2C uses only two wires, additional complexity is added in handling the overhead of addressing and data acknowledgement
  I2C can be inefficient when simple configurations and direct linking can be interfaced

## Slide 37: USB Interface
Future Technology Devices International, LTD.
VDIP1 Module
Utilizes FTDI’s VN1CL USB Host Controller IC
Handles USB protocol
Supports SPI interface with PIC18F
USB A Type Socket
$24.50

## Slide 38: Data Output Schematic

## Slide 39: Microchip’s MDD
Memory Disk Drive (MDD) Library
  Free
  Wide range of support
  Provides method of interfacing files and directories
  FAT12, FAT16, and FAT32
  Most popular with SD cards and USB thumb drives

## Slide 40: File Format
Input (TXT file)
mode
frequency
time0, force1,force2, force3, force4, force5, force6, force7, force8
time1, force1,force2, force3, force4, force5, force6, force7, force8
Output (CSV file)
mode,AL/LT
frequency,00,Hz
time,Channel1,Channel2,Channel3,Channel4,Channel5,Channel6,Channel7,Channel8
00:00:00:00:00,00.0,00.0,00.0,00.0,00.0,00.0,00.0,00.0

## Slide 41: Graphic User Interface

## Slide 42: GUI
Provide an interface for the user to control the MTEC on the touch screen
Display data and progress while MTEC running
Programmed in C
Graphics.h library provides functions to draw graphics on screen
Touch simulated using mouse-click functions in C

## Slide 43: Instruction Input GUI

## Slide 44: Instruction Input GUI II

## Slide 45: Instruction Input GUI III

## Slide 46: Administrative

## Slide 47
Chart: Completion Summary
| Testing | Programming | Parts Acquisition | Design | Research
 | 0.1 | 0.3 | 0.5 | 0.85 | 0.9

## Slide 48: Budget
Chart: 
 | 881.13 | 348.65
Component | Price | Qty | Projected | Actual | Acquired
MTEC Components |  |  |  |  | 
PIC18F Dev Kit | $165.00 | 1 | $103.00 | $165.00 | Y
GLCD w/ Touch Screen | $87.00 | 1 | $61.56 | $87.00 | N
USB VDIP1 Module | $24.50 | 1 | $23.42 | $24.50 | N
SD Card Socket | $9.95 | 1 | $9.95 | $9.95 | N
Breakboard | $9.95 | - | - | - | N
Pactec Enclosure | $28.20 | 1 | $28.20 | $28.20 | N
PIC18F87K22 Plugin Module | $25.00 | 1 | - | $25.00 | Y
Sub Total |  |  | $226.13 | $314.65 | 
Rig Components |  |  |  |  | 
Futek LCM 300 FSH02632 Load Cell | $450.00 | 2 | $575.00 | $900.00 | Y
Linear Actuator Firgelli L12-50-210-06-I | $80.00 | 1 | $80.00 | $80.00 | Y
Transducer LD621-15 | $455 | 2 | - | $910.00 | N
Sub Total |  |  | $1230.00 | $1890.00 | 
Grand Total |  |  | $1456.13 | $2204.65 | 

## Slide 49
Challenges
Acomodating the response time of the actuators.
Analog signal alterations created when modifying the sensor signals.
Programming the GUI of the GLCD.
Parallel implementation of the SD and USB.
Incorporating a surface mounted microcontroller.