## Slide 1
Presenters:
Brandon Sbert (EE)
Raj Bose (EE)
Bianca Belmont (CPE)
Ricardo Wheeler (EE)
Blood Pressure Tester
Sponsors:
Texas Instruments
Workforce Central Florida
Mentor:
Herb Gingold (TI)

## Slide 2: Project Description
Build an Automatic Blood Pressure Tester utilizing the Oscillometric Method (indirect)
Low Power
Wireless Display

## Slide 3: Goals and Objectives
To be worn on upper arm
Battery powered
Simple user operation (one button device)
Integrate safe procedures into design
Implement wireless component
Calculate Blood Pressure reading (SYS DIA)
Transmit results wirelessly to display
Receive data from wireless module
Display Blood Pressure data
Error detection

## Slide 4: Specifications
Power Supply 4 AAA rechargeable batteries (3v)
Power Life is 60 BP runs
Automatic using Micro motor (6V) / Micro Valve (6V)
Oscillometric
Accuracy of sensor plus or minus 3mmHg
Pressure range of 20mmHg to 280mmHg (cuff)
Adjustable cuff
Wireless range 1m <range> 2m
Display 138X110 grayscale, dot-matrix LCD

## Slide 5: Hardware Block Diagram

## Slide 6
BP = SYS (high pressure contracting) / DIA (low pressure relaxed)
Blood Pressure Monitoring

## Slide 7: Blood Pressure Monitoring
Many invasive and non invasive methods exist
Similarity of 3 non invasive methods
  all 3 use an occlusion cuff
  all 3 record pressure values upon the turbulent re-entry of
  blood to lower arm
  all 3 inflate cuff to about 30 mmHg above average systolic
  pressure to cut off blood flow to the lower arm
  Palpitation – touch – direct method
  Auscultatory – hearing – direct method
  Oscillometric – algorithmic – non direct
BP = SYS (high pressure contracting) / DIA (low pressure relaxed)

## Slide 8: Auscultatory Method
Direct Method
Based on 5 auditory events
(sound / silence)
Heard with stethoscope or microphone
Record meter pressure at first and last event to obtain SYStolic and DIAstolic pressure values
BP = SYS / DIA

## Slide 9: Auscultatory Method

## Slide 10: Oscillometric Method
BP = SYS (high pressure contracting) / DIA (low pressure relaxed)
Utilized in our device
Indirect Method
Cuff wall assumed one with the skin
  Movement of skin due to turbulent blood flow pulses upon re – entry
  Creates air turbulence in cuff
Algorithm uses two sets of data:
  Originating from a mixed signal obtained by a pressure sensor connected
  to an occlusion cuff
  Calculates a systolic pressure and diastolic pressure for a blood pressure reading

## Slide 11: Oscillometric Method

## Slide 12: Oscillometric Method
  Data set 1
    Cuff pressure vs. time
  Data Set 2
    Only MAP Mean Arterial Pressure obtained from signal
      Average arterial pressure during one heart cycle
      MAP = DIA + 1/3 (SYS – DIA)
    MAP Mean Arterial Pressure PEAK amplitude of signal
      Counterintuitive: MAP is the PEAK of a signal of re-entry pulses
      SYStolic pressure is assumed to be the highest pressure in the heart cycle
      SYStolic and DIAstolic points in time in relation to MAP

## Slide 13: Mechanical

## Slide 14: General Picture of the Mechanical Parts

## Slide 15: Motor
Model: P54A02R
Cylinders: 3
Rated Voltage: DC 6V
Flow (No Load): 1.8L/min
Current (No Load): 170mA
Max Current: 290mA
Max Pressure: 95kPa
Noise: 50dB

## Slide 16: Cuff
Model: D-Ring
Upper Arm
Standard adult cuff which has a circumference between 9-13 inches
Used for home-monitoring and self-application environments
It provides great flexibility, and it is light

## Slide 17: Solenoid Valve
Model: KSV05B
Rated voltage: DC 6V
Rated Current: 60mA/45mA
Exhaust time: Max. 6.0 seconds from 300mmHg reduce to 15 mmHg at 500CC tank
Leakage: Max. 3mmHg/min from 300mmHg at 500CC tank.

## Slide 18: Mechanical Valve
Maintains a slower linear deflation rate
Optimal for pressure sensor sampling:
160 – 80 mmHg (Cuff Pressure)

## Slide 19: Pressure Sensor
Freescale MP3V5050GP
Internal amplification
Low pass output to avoid noise
Required
  7mA constant current input
  3.3 V input
Input Range 0 - 50 kPA ( 0 - 7.25psi)
Output Range 0.06 – 2.82 V out
Transfer Function
Vout = Vin * (0.018 * kPa + 0.04)
7.50061683 mmHg per 54 mV
BP = SYS / DIA = mmHg

## Slide 20: Pressure Signal

## Slide 21: Pressure Signal

## Slide 22: Oscillation Signal
Systolic
  Point in time when signal is 55% of the MAP amplitude
Diastolic
  When signal has decreased by 85% of MAP amplitude

## Slide 23: MCU

## Slide 24: MSP430F5438A
MSP430F5438A Features:
  16-bit Ultra-low power microcontroller
  256KB Flash
  16KB RAM
  High performance 12-bit analog-to-digital (A/D) converter
  Real-time clock module
Language: C
Implementation: Code Composer Studio v5.1
Schematics: TINA and WEBENCH Designer

## Slide 25: Software Diagram

## Slide 26: WEBENCH

## Slide 27: Hardware Block Diagram

## Slide 28: Wireless

## Slide 29: Wireless Options
Data | CC1101 EM – Sub 1GHz radio | XBee 1mW Chip Antenna - Series 1 (802.15.4)
Power | 3.3V @34.2 mA | 3.3V @45 mA
Frequency | 868-915 MHz | 2.4 GHz
Protocol | RF | RF
Range | Short Range | Short Range
Antenna | Wire | Chip
Support | Little | A lot

## Slide 30: XBee 1mW Chip Antenna - Series 1 (802.15.4)
Protocol: RF
Frequency: 2.4 GHz
Power: 3.3V @ 45mA
Range: 300ft (100m) range
Antenna: Chip Antenna

## Slide 31: Wireless Block Diagram

## Slide 32: Wireless Design
Pin1: Vin at 3.3V
Pin2: Dout Connected to RX
Pin3: Din Connected to Tx
Pin10: Ground
Pin5: RESET
Pin9: Digital Input/Sleep Control
Pin12: Clear-to-send flow control
Pin13: Module Status Indicator
Pin16: Request-to-send flow control

## Slide 33: Power Source

## Slide 34: Battery
4x AAA batteries: 6V
Alkaline Batteries
Power life 60 BP runs

## Slide 35: Power Regulator for the Motor/Valve
Model: LM3488
Efficiency: 80%
Switching Frequency (Max): 1000kHz
Switching Frequency (Min): 100kHz
Vin (Min): 2.95V
Vin (Max): 40V
Vout: 2.97V to 40V
It will be supplying the motor and the valve

## Slide 36: Schematic of the Power Regulator for the Motor/Valve

## Slide 37: Power Regulator for the MCU/PS/Wireless
Model: TPS62122
Efficiency: 96%
Vin (Min): 2V
Vin (Max): 15V
Vout (Min): 1.2V
Vout (Max): 5.5V
It will supply the MCU, Pressure Sensor and Wireless

## Slide 38: Schematic of the Power Regulator for the MCU/PS/Wireless

## Slide 39: TS12A4514 - Switch
Single pole/single throw (SPST), low-voltage, single-supply CMOS analog
It is normally open (NO)
These CMOS switches can operate continuously with a single supply between 2 V and 12 V
Will be turned ON and OFF by the MSP430F5438 which will be sending 3.3 V (High)

## Slide 40: TPS1101PWR MOSFET
The TPS1101 is a single, P-channel, enhancement-mode MOSFET
It is a normally open (NO)
It is the ideal high-side switch for low-voltage, portable battery-management systems where maximizing battery life is a primary concern
It will supply enough current of 290mA to the motor
It operates under 6 V

## Slide 41: Printed Circuit Board (PCB)

## Slide 42: PCB Design
Data | Values
Software | EAGLE
Manufacturer's | PentaLogix, Inc. & ”Just In Time”
Size | 4” x 3”
Layers | 2
Components | 62

## Slide 43: Eagle Designed PCB (Both Layers)
Dimensions are in mm

## Slide 44: Bill OF Materials (BOM)

## Slide 45: PCB Provided by PentaLogix

## Slide 46: PCB Assembled by “Just In Time”

## Slide 47: Testing
Test Runs
135/87
131/85
134/86
130/84
Brandon’s Test Runs From Project
Brandon’s Test Run at Publix

## Slide 48: Work Distribution
Component | Brandon Sbert | Bianca Belmont | A. Raj Bose | Ricardo Wheeler
MCU/Coding | 5% | 80% | 10% | 5%
Power Design | 10% | 5% | 5% | 80%
Filter Design | 5% | 5% | 80% | 10%
Wireless Design | 80% | 10% | 5% | 5%
Schematic/PCB | 80% | 5% | 10% | 5%

## Slide 49: WCF Budget
Component | Quantity | Total Price
Batteries | 8 | $20.00
BP Motor | 2 | $10.00
BP Pump | 3 | $30.00
BP Valve | 3 | $9.00
BP Cuff | 2 | $40.00
MCU | 3 | $3.00
Op-Amps | 5 | $10.00
Resistors | 10 | $7.00
Capacitors | 10 | $10.00
EXP Board/Display | 1 | $200.00
Pressure Sensor | 4 | $65.00
Wireless EM | 2 | $240.00
PCB Board | 1 | $55.00
Sub Total: | 54 | $699.00

## Slide 50: Actual Budget
Component | Quantity | Total Price
Batteries | 20 | $15.00
BP Motor/pump | 2 | $25.00
BP Valve | 2 | $25.00
BP Cuff | 2 | $25.00
Components | 62 | $120.00
Tina Software | 1 | $89.00
PCB | 2 | $225
MSP430F5438 | 6 | FREE
Xbee Module | 2 | $50
MSP430F5438 Experimenter Board | 4 | FREE
CC1101DK868-915 | 1 | FREE
MSP-FET430U5X100 | 1 | FREE
Shipping | N/A | $100
Sub Total | 105 | $674

## Slide 51: Problems
Testing LM Regulator
LM Regulator not working on original PCB

## Slide 52: Questions?