## Slide 1: Direct Absorbance Bilirubin Spectrometer (DABS)
Bernardin Dezius, CpE
Juan Gonzalez, EE
Kevin Kuzius, EE
Kevin Landau, PSE
Andrea Wetteland, PSE
Group 4

## Slide 2: Motivation
Jaundice afflicts up to 60% of all newborns.
Build-up of bilirubin in the blood that the newborn's liver can't efficiently filter
Causes yellowing of eyes and skin
Most cases are mild and can be treated with exposure to sunlight.
Higher levels can cause brain damage if not treated.

## Slide 3: Goals & Objectives
The goal of this project is to provide a quick, safe, and cost-effective medical tool in determining high levels of bilirubin in newborns.
DABS device is for clinical use that can quickly determine the severity of jaundice based on the absorbance of bilirubin in a blood sample within seconds.

## Slide 4: Executive Summary
•Bilirubin and hemoglobin contribute to the absorption measured with a blue wavelength.
•Only hemoglobin contributes to the absorption at 540 nm.
•Based on hemoglobin and oxygen levels in a normal, full-term neonate, the hemoglobin absorption at 540 nm can determine the amount of absorption is due to hemoglobin at 460 nm.

## Slide 5: Specifications

## Slide 6: Block Diagram

## Slide 7: Proxy Samples
Curcumin
Bilirubin
Betalain
Hemoglobin

## Slide 8: Calculating Concentration
• Beer-Lambert Law
A = lcε = -log(I0/IT)
• Comparison of incident (blank) to transmitted
• Hold separate curcumin and betalain concentrations constant to find ε for each wavelength.
• Use ε coefficients to calculate unknown concentrations

## Slide 9: Sample Preparation
Concentrations must be consistent
Powdered beetroot and turmeric dissolved in methanol for greatest accuracy
Time constraints include photosensitivity of dies and sediment

## Slide 10: Light Sources
Requirements
•Power requirements: ~4 V
•Power output: 5-50 mW
•Wavelength output: ~460 nm, ~550 nm, with low FWHM
•Low-cost
Laser Diodes vs LEDs
•LEDs cheaper, more efficient
•Laser diodes have much narrower FWHM range

## Slide 11: FWHM Spectral Emissions
•A wide spectral width makes it almost impossible to attribute amounts of absorption to specific wavelength ranges.
•Equal and specific spectral widths for each wavelength are necessary to accurately determine absorption

## Slide 12: Lasers – Temperature Requirements
•The power output be same for the blank and the transmitted readings
•Reading at 1 sec before slope begins to increase
•Purpose of fan to mitigate this temperature increase

## Slide 13: Detector
Convert optical signal to an electrical signal
High sensitivity
Spectral range that includes 460 nm and 550 nm
Linear response
Use Beer-Lambert Law to calculate the concentration of the desired substance

## Slide 14: Detector Comparison
Size Comparison

## Slide 15: Detector – BPW21R
•This detector had the highest sensitivity (9 nA/lx) and a narrow spectral range
•Linear response
•Ease of use for testing and implementation into the circuit
•Larger radiant sensitive area 7.5 mm2 as opposed to 7.02 mm2

## Slide 16: Proof of Concept

## Slide 17: Microcontroller Technology Comparison
Parameter | MSP430FR6989 | MSP432P4111 | ARM Cortex-M3
GPIO | 63 programmable (80pin) 83 programmable (100pin) | 84 programmable | 54 Digital 12 Analog Input 2 Analog Output
ADC Resolution | 12 Bit | 16 Bit | 12 Bit
Clocking Features | 16 MHz (max) | 48 MHz (max) | 84 MHz (max)
Memory (Flash) | 128 KB | 2048 KB | 512 KB

## Slide 18: Microcontroller
Coordinates all operations
Stops modules from conflicting and wasting power
Low-power modes
Actual Size: 14mm x 14mm
Texas Instruments
Notes: CPU, MCLK, DCO, SMCLK, ACLK

## Slide 19
User Interface
BOOSTXL-K350QVG-S1

## Slide 20: Display’s consumption
Internal circuit, turns off display
Touchscreen - sampling driven by touch interrupt

## Slide 21: Flowchart

## Slide 22: Use Case Diagram

## Slide 23: Touchscreen – Capacitive vs. Resistive
Resistive
Capacitive
Multi-touch
Unusable with most gloves
Stylus incompatible
Sensitive to touch
More expensive
Responsiveness can vary per implementation in various products
Non multi-touch
Usable with gloves
Stylus compatible
Less sensitive to touch
Less expensive
Lacking in display quality
Notes: We picked resistive over capacitive because of out-of-the-box compatibility and low cost
The biggest differences being touch sensitivity, touch versatility, and multi-touch capability
Capacitive would have been nice option if we wanted the include multi-touch features

## Slide 24: Power Input
Lithium-Ion Battery:
4.2 V Max @ full charge / 3.7 V nominal
2500 mAh capacity (10 Wh)
Adafruit - $14-15
USB Connection:
5 V Output
Will be used to charge the battery
Will supply power to the system when the battery is charging

## Slide 25: Battery Charging- TI’s MCP73831

## Slide 26: Other Components
Brushless Cooling Fan (x2)
Rated Voltage of 3.3 or 5 V
13000 RPM
180 mA
Amazon - $8
Temperature Sensor (TI LM35)
4 – 30 V
Range of -55°C to 150°C
< 60 uA current draw

## Slide 27: Battery Switching
USB Connection
(5V)
Battery
System

## Slide 28: PCB Schematic
MCU-MSP430FR6989

## Slide 29: PCB Schematic
LCD I/O

## Slide 30: PCB Schematic
Emulator I/O

## Slide 31: PCB Schematic
Cooling System
I/O

## Slide 32: PCB Schematic
Optical Connections
Laser (2)
Photo Detector

## Slide 33: PCB Schematic
Amplifier Circuit

## Slide 34: PCB Schematic
DC-DC Converters

## Slide 35: PCB Schematic
Battery Charging Circuit

## Slide 36: PCB Layout

## Slide 37: Optical Housing
Components: Lasers, cuvette holder, temperature sensor and detector
Lasers will be placed at an angle and must cross the front surface of the cuvette and reach the detector
All measurements are in inches
Inside view w/inner box dimensions
Top view w/inner and outer box dimensions

## Slide 38: Outer Housing
Contain the PCB and other electronics such as the display and fans
Housing will be made using acrylic plastic which is lightweight, durable and easily machinable
Top view w/inner and outer box dimensions​
Lid of the housing w/respective measurements
All measurements are in inches​

## Slide 39
Administrative Content

## Slide 40: Project Budget
System | Part | Budgeted | Actual
Optical | Blue Laser Diode | $20 | $20
 | Green Laser Diode | $20 | $16
 | Cuvettes | $20 | $14
 | Photodetector | $50 | $15
 | OpAmp OPA170 | $5 | $0*
Power | Battery & Parts | $50 | $14
MCU and Screen | MSP430FR6989 80 Pin | $30 | $3
 | LCD Display | $100 | $26
PCB | Board | $50 | $40
 | BOM (Components) | $50 | $47
Temperature Control | Fan | $10 | $10
 | Temperature Sensor | $15 | $0*
Housing | Acrylic Plastic + Machine Work | $55 | $0*
Total |  | $475 | $205

## Slide 41: Responsibilities
 | Andrea | Kevin K | Juan | Bernardin | Kevin L
Power |  | S | P |  | 
PCB |  | S | P |  | 
Light Source | P |  |  |  | S
Optical Alignment | P |  |  |  | P
Detector | S |  |  |  | P
Amplifier |  | P | S |  | 
MCU/Software |  | S |  | P | 
Display/ Touchscreen |  | S |  | P | 
Housing | S |  |  |  | P

## Slide 42: Progress

## Slide 43
Questions?