## Slide 1: High Precision Temperature Controller
Group 13
Ashley Desiongco
Stacy Glass
Martin Trang
Cara Waterbury

## Slide 2: Objectives
Replace COTS controller
  More Efficient
  More Economical
Use modern technology
  Part selection must consider production life
Notes: 10 year prod cycle

## Slide 3: Application
Extended Area
Will use 2 Type T T/C or 4 RTDs
From -30°C to 700°C
Cavity
Will use 2 Type S T/C
From 50°C to 1200°C
Notes: Currently uses single RTD, one Type S

## Slide 4: Top Level Block Diagram

## Slide 5: Analog Subsystem

## Slide 6: Sensor & Reading Specifications
Must stabilize within +/- .5°C
Read a minimum of:
  2 differential thermocouple signals
  5 RTD signals
Convert to digital signal and send to PIC
All noise/drift must be accounted for

## Slide 7: Sensor Types
Thermocouples
Type S
  20 ⁰C min
  1300 ⁰C max
  0.1107 mV to 13.17 mV
  Cavity source
Type T
  -30 ⁰C min
  400 ⁰C max
  -1.21 mV to 20.87 mV
  Extended area source
RTDs
PT100
  -30 ⁰C min
  400 ⁰C max
  Extended area source:
    88.22 Ω to 247.09 Ω
  Cold junction comp:
    100 Ω to 123.24 Ω

## Slide 8: Block Diagram

## Slide 9: Differential Op Amp
Differential output conditioning Op Amp
VOCM = 2.5 V reference voltage
Internal precision 10kΩ resistors

## Slide 11: RTD Readings
RTD ladder
Requires only 1 precision resistor
Must match min input requirements of AD converter

## Slide 12: Schematic

## Slide 13: A-D Converters
AD7797
24 bit resolution
1 differential input
SPI interface
Internal gain amplifier fixed at 128
Used for heater (TC) reading
AD7718
24 bit resolution
8 channel input MUX
SPI interface
Internal PGA of 1 to 128
Used for all RTD readings and secondary TC reading

## Slide 14: Reference Voltage Considerations
Component | Current Draw
AD7797 | 1 μA
AD7718 | 1.25 μA
AD8476 – Op Amp (2) | 5 μA
RTD Ladder | 713 μA
TOTAL | 720.25 μA
Vout = 2.5 V
Iout = 40 mA
Temp drift = 3ppm/ ⁰C

## Slide 15: Microcontroller
Notes: Once the information is on the ADC, it’s the MCU job to process that information

## Slide 16: Microcontroller Specifications
Capable of Communicating with 8 Peripheral Devices.
Capable of Handling RS-232, RS-485, USB, and Ethernet Protocols.
Capable of performing signed, floating point math.
Notes: These are the main three specs we have for it. First

## Slide 17: PIC32MX150F128B
2 SPI Interfaces
2 UART Interfaces
Full-featured ANSI-Compliant C Programming Language
Notes: And the MCU that fit the bill was the pIC32mx150F128B

## Slide 18: General Design
Two PIC32MX150F128B connected in Master-Slave configuration.
Slaves will be customized to serve a single purpose.
Master will handle outside communication and slave coordination.
Notes: Our general configuration for the microcontroller is to have two pIC32MX150128B connected in Master-Slave Configuration

## Slide 19: Pinout

## Slide 20: Peripherals (from the Master)
MAX232 – RS232 – UART
MAX481 – RS485 – UART
MCP2200 – USB – UART
ENC28J60 – Ethernet – SPI
µLCD-32032 – Display – UART
PIC32MX150F128B – Slave – SPI
Independent 8-level deep FIFO TX/RX UART Buffers
Independent 4-level deep FIFO TX/RX SPI Buffers onboard the PIC32MX150F128B

## Slide 21: Development Environment
MPLABX using MPLAB C32
Simulation Capability
Debugging
  Using PICKIT3

## Slide 22: Display

## Slide 23: Requirements
Touch Screen
Low-Cost
Fit in existing chassis
Interface easily to microcontroller

## Slide 24: 4D-Systems uLCD32 (GFX)
Built in Graphics
Controller
Easy 5-pin interface
On-board Audio
Micro-SD card connector
Expansion Ports
Built in Graphics Libraries

## Slide 25: Features
2
3
4
5
6
1
480x272 Resolution
Expansion Ports (2)
5 Pin Serial Programming Interface
PICASO-GFX2 Processor
Micro-SD Card Slot
1.2W Audio Amplifier with Speaker
3.2”

## Slide 26: Hardware Interface
Easy 5 pin interface
  Vin, TX, RX, GND, RESET
Also used to program display with 4D Programming Cable

## Slide 27: PICASO-GFX2 Processor
Custom Graphics Controller
Configuration available as a PmmC (Personality-module-micro-Code)
PmmC file contains all low level micro-code information

## Slide 28: Audio/Micro-SD Card
Audio support is supplied by the PICASO-GFX2 processor, an onboard audio amplifier and 8-ohm speaker
Executed by a simple instruction
Micro-SD card is used for all mulitmedia file retrieval
Can also be used as general purpose storage

## Slide 29
Temperature displayed at all times
Change current set point option

## Slide 30: Power

## Slide 31: Power
Part | Current (mA) | Voltage (V) | Quantity | Power (mW)
ADC | 0.65 | 5 | 1 | 3.25
ADC | 0.325 | 5 | 1 | 1.625
ADC | 0.65 | 3.3 | 1 | 2.145
ADC | 0.325 | 3.3 | 1 | 1.0725
OpAmp | 0.33 | 5 | 2 | 3.3
Ref | 0.8 | 5 | 1 | 4
Quad Buffer | 30 | 5 | 1 | 150
RS485 | 0.9 | 5 | 1 | 4.5
RS232 | 15 | 5 | 1 | 75
USB | 95 | 5 | 1 | 475
Ethernet Controller | 180 | 3.3 | 1 | 594
Display | 150 | 5 | 1 | 750
Microcontroller | 50 | 3.3 | 2 | 330
4:1 MUX | 75 | 3.3 | 1 | 247.5
TOTALS | 649.31 |  |  | 2641.393

## Slide 32: Power Block Diagram
LS25-5
90 – 240 Vac
5V
ADC RS485
OpAmp RS232
Ref. Display
Buffer USB
LT1129-3.3
Ethernet
Microcontroller
4:1 MUX
ADC
3.3V

## Slide 33: Temperature control Method

## Slide 34: PID Requirements
Eliminate noise
Minimize overshoot
More efficient than standard PID

## Slide 35: Nested PID
Influence of parameters:
  P = Decreases rise time
  I = Eliminates SS Error
  D = Decreases overshoot and settling time
Initial loop encompasses entire temperature range using only P and D parameters
Next loop focuses on a smaller range and uses P, I and D

## Slide 36: Analog system software design

## Slide 37: Interfacing with AD7797
Thermocouple Reading
Initialize AD7797 to the following settings:
  Unipolar Mode: 0 – 20 mV
  Sampling Frequency: 123 Hz
  Clock Source: Internal 64 kHz
  Converting Mode: Continuous Conversion Mode
Reading data output register:
  Send 0x58FFFFFF to DIN of AD7797 – Single Read Operation

## Slide 38: Interfacing with AD7718
CJC Reading
Initialize AD7718 to the following settings:
  Unipolar Mode
  Programmable Gain: 128
  Sampling Frequency: 105.3 Hz
  Chopper Enabled
  Converting Mode: Continuous Conversion Mode
  Channel Select: AIN(+) – AIN3; AIN(-) – AIN4
Reading data output register:
  Send 0x44FFFFFF to DIN of AD7718 – Single Read Operation

## Slide 39: Temperature Conversion
Acquire CJC equivalent voltage reading
Acquire thermocouple voltage
Subtract CJC voltage from thermocouple voltage
Translate to temperature using NIST Standard Tables.
AD7718 Formula
AD7797 Formula

## Slide 40: Peripheral software design

## Slide 41: General Overview
No Interrupt Driven Events
  Constant Polling Transmit/Receive Buffers for SPI and UART
Master PIC handles data transfer to and from the Display and Slave PIC
Master PIC serves as a slave to the Computer Interface.
Custom LABVIEW software to handle all computer interfacing.

## Slide 42: Display software Design

## Slide 43: General Overview
Polls RX buffer for command from master
  0x01: master to send current temperature
  0x02: master to send new set point
  0x03: master requests new set point from display
Handles touch events
  Uses internal functions to determine location of touch events

## Slide 44: Software Tools
1. 4D Workshop IDE
2. PmmC Loader
3. Graphics Composer
4. FONT Tool

## Slide 45: Temperature Formatting
Data sent in 3 bytes from master or display
  Display UART is limited to 1 byte
First Byte: Contains tenths place (upper four bits) and ones place (lower four bits)
Second Byte: Contains tens place (upper four bits) and hundreds place (lower four bits)
Third Byte: Contains Thousands place (upper four bits) and sign/check bit (lower four bits)
  Fourth bit must be set high for data to be valid.

## Slide 46: PID software Design

## Slide 47: General Overview
Compare Set Point temperature with Current temperature
Check if the current temperature is within the proportional band
Accumulate error (for Integral Action) and store previous temperature (for Derivative Action)
Calculate Proportional, Integral, and Derivative terms
Translate PID terms into varying duty cycles for PWM output

## Slide 48: Testing

## Slide 49
Testing OpAmp
Testing AD7797 (via PIC32 Starter Kit)
Testing AD7797 (via PIC32MX150F128B)
Full System Integration Testing

## Slide 52: Pid parameter Testing

## Slide 53: Trial 1
P Band = 5%
Repeats per Minute= .65
Derivative Time= .001
Set Point = 600.0°C

## Slide 54: Trial 2
P Band = 5%
Repeats per Minute= .50
Derivative Time= .01
Set Point = 600.0°C

## Slide 55: Trial 3
P Band = 5%
Repeats per Minute= .50
Derivative Time= .01
Set Point = 700.0°C

## Slide 56: Work Breakdown
 | Ashley | Martin | Cara | Stacy
Analog Hardware | 95% | 5% | - | -
Digital Hardware | - | 80% | - | 20%
Display | - | 5% | 95% | -
Software | 5% | 10% | 5% | 80%
Power | - | - | 100% | -

## Slide 57: Budget
Parts | 
Digital Devices | $ 21
Analog Devices | $ 30
Passive Devices | $ 62
Power Devices | $ 20
Display | $ 101
Board Fabrication | $ 80
Programming Tools | $ 52
TOTAL | $ 366
Goal: $500

## Slide 58: Educational Experience
Conflicting Reprogrammable pin assignment definitions
LATx versus PORTx
Three Tier SPI handshaking
Board Population

## Slide 59: QUESTIONS?